|  | Department: Respiratory Services | Date Originated: April 2010 Date Reviewed/Revised: |
| --- | --- | --- |
| PROCEDURE | Topic: Critical Care – Esophageal Pressure Monitoring With the AVEA Ventilator Number: RTD5195 | Related Links: |
| PRINTED copies of Clinical Practice Standards and Procedures may not be the most recent version. The OFFICIAL version is available on the PHC Intranet – Policies & Manuals web page. APPLICABLE SITES: St. Paul’s Hospital GENERAL INFORMATION: Lung protective strategies used during mechanical ventilation should ideally provide sufficient transpulmonary pressure (airway pressure – pleural pressure) to maintain gas exchange while minimizing repetitive alveolar collapse and overdistention, and volutrauma. This is generally gauged by monitoring gas exchange via ABG and airway/plateau pressures. However for this to be true, we are also assuming that chest wall compliance is normal, and as such has a minimal impact on overall respiratory system compliance. In patients where chest wall compliance is reduced (i.e. obesity, ascites, bowel edema), elevated airway/plateau pressures may not accurately reflect alveolar overdistention. In fact, in these types of cases a strict adherence to the aforementioned lung protective strategies may result in alveolar collapse, poor gas exchange, and ongoing ventilator induced lung injury. Using an esophageal pressure measurement to estimate pleural pressure will allow for better estimation of transpulmonary pressures. In patients with reduced chest wall compliance, transpulmonary plateau pressure is a better reflection of alveolar distension, while transpulmonary PEEP is a better estimation of alveolar recruitment. CAUTIONS & CONTRAINDICATIONS: Esophageal varices Esophageal trauma Recent esophageal surgery Coagulopathy Platelets < 80 INR > 2 Post heart or lung transplant Elevated intracranial pressure DEFINITIONS: Transpulmonary Plateau Pressure (Ptp Plat): Ptp Plat = Pplat – Pes where Pes = esophageal pressure measured during inspiratory hold and estimates pleural pressure Target Ptp Plat ≤ 25 cmH20 Transpulmonary PEEP (Ptp PEEP): Ptp PEEP = PEEPtotal - Pes where Pes = esophageal pressure measured during expiratory hold and estimates pleural pressure Target Ptp PEEP ≥ 0 cmH2O see also Ptp PEEP-guided FiO2 Table noted below For Viasys Avea EQUIPMENT: AVEA ventilator SmartCath 8Fr Adult, Esophageal Pressure Monitoring Tube Set SmartCath Extension, Esophageal Pressure Monitoring Tube Set Water soluble lubricant Magill forceps ( optional) Laryngoscope ( optional) PROCEDURE FOR PREPARATION OF THE ESOPHAGEAL CATHETER AND VENTILATOR: A balloon test must be performed prior to placement of the esophageal balloon. Connect the SmartCath esophageal extension tubing to the Pes port. Ensure the pins align properly with the Pes port. Once connected, slide the sleeve of the extension tubing connector over the Pes port mount. Connect the esophageal catheter to the extension tubing. Ensure the pins are aligned for a proper fit. Once connected slide the sleeve of the catheter connector over the extension tube port mount. Press SCREENS to access the Screens page. Press MANEUVERS to access the Maneuvers page. Press ESOPHAGEAL to access the Esophageal Maneuver Settings page. NOTE: On the Esophageal screen, a message indicating the balloon status appears in the message bar. If BALLOON DISCONNECTED appears, reconnect the tubings. A: If using 6 french catheter, Select ADULT SIZE ESOPHAGEAL BALLOON and press ACCEPT. B: If using 16 french combo tube, Select Adult Size Nasogastric Esophageal Balloon and press accept. Verify correct wording on the selection menu The Balloon Test icon appears in the Maneuver Settings screen. Press BALLOON TEST. The alert message REMOVE BALLOON FROM PATIENT PRIOR TO LEAK TEST appears. Press CONTINUE. The ventilator will now inflate and deflate twice to confirm the integrity. Following this test, the message ADULT BALLOON TEST PASSED will appear in the message bar and a Pes software key becomes available. NOTE: Pressing the Pes OFF key will inflate the esophageal balloon and the software indicator will change to Pes ON. NOTE: Do NOT turn Pes ON until the catheter is in the patient and measurements are ready to be made. PROCEDURE FOR PLACEMENT OF THE ESOPHAGEAL CATHETER: Measure the distance from the tip of the Nose (N) via the Earlobe (E) to the Xiphoid process (X). This is referred to as the NEX measurement. NOTE: For oral insertion measure from the incisors instead of the nose. Make note of the NEX measurement in centimeters, and use that value to determine the insertion depth. NOTE: A measurement of 40 cm (third hatch mark) will place the esophageal balloon in the lower ⅓ of the esophagus for most adults. The distance between hatch marks on the catheter is 5 cm, with the first hatch mark located 30 cm from the catheter tip. With the patient in a semi-Fowlers position, insert the catheter in short strokes 2-3 cm at a time. This will help to prevent coiling in the pharynx. The catheter should insert with minimal resistance. For oral insertion, try to follow the path of the endotracheal tube through the oral cavity. NOTE: The RN, intensivist, fellow, or resident will be responsible for insertion of the esophageal balloon catheter. Query add RRT? Once the catheter has been inserted, access the Esophageal Maneuver Settings screen and turn Pes ON. The balloon will now be automatically inflated over a duration of 5 – 10 seconds. Verify catheter placement by observing for cardiac oscillations on the Pes waveform, which will appear as a saw-tooth pattern. NOTE: In the non-spontaneously breathing patient, both the Paw and Pes should rise in the same direction with each mechanically delivered breath. Placement verification via belly pump during expiratory hold Too high will show positive deflection at the Ptp Too low ( in the gut) will show negative deflection at the PTP Good placement will show minimal deflection at PTP ( no more than 1-2 cmh20) with a distinct cardiac oscillation Balloon catheter in the lung will show minimal deflection at the PTP with a distinct lack of cardiac oscillation NOTE: If the lungs are very non-compliant, the Pes scalar may need to be adjusted to visualize the cardiac oscillations. Note the markings on the catheter and record the position in centimeters on the Critical Care Respiratory Therapy Flowsheet. Final verification can be achieved with a chest x-ray. The NG combo tube has a distinct 1 cm radiopaque line at the distal tip and another at the distal end of balloon to help identify placement. PROCEDURE FOR OBTAINING PRESSURE MEASUREMENTS: Ensure patient is supine with HOB 30 degrees prior to doing measurements. Screen Display Setup: Once the esophageal balloon has been inserted and correct position has been verified, set up the MAIN SCREEN as follows: Waveforms Display: Paw (cmH2O) Pes (cmH2O) Ptp (cmH2O) Monitored Values Display: Pplat (cmH2O) Ptp Plat (cmH2O) Ptp PEEP (cmH2O) Transpulmonary Plateau Pressure (Ptp Plat): Press and hold the Inspiratory Hold key. Press Freeze and scroll to the point where the inspiratory hold was performed. **Target Ptp Plat ≤ 25 cmH2O (maintain low Vt strategy)** Transpulmonary PEEP (Ptp PEEP): Press and hold the Expiratory Hold key. Press Freeze and scroll to the point where the expiratory hold was performed. **Target Ptp PEEP ≥ 0 cmH2O** or **Set ventilator PEEP to target Ptp PEEP at the level below based on the FiO2** Ptp PEEP-guided FiO2 Table FiO2 0.4 0.5 0.5 0.6 0.6 0.7 0.7 0.8 0.8 0.9 0.9 1.0 Ptp PEEP 0 0 2 2 4 4 6 6 8 8 10 10 PROCEDURE FOR PLACING ESOPHAGEAL PRESSURE MONITORING IN STANDBY (for patient transport): Access the Esophageal Maneuver Settings screen and turn Pes OFF. This will deflate the balloon. Disconnect the SmartCath esophageal pressure monitoring catheter where it connects with the SmartCath extension tubing. Do NOT disconnect the SmartCath extension tubing from the ventilator. If the extension is disconnected from the AVEA, the balloon test will need to be repeated and the catheter will have to be removed from the patient. After transport, reconnect the SmartCath esophageal pressure monitoring catheter to the extension tubing. Access the Esophageal Maneuver Settings screen and turn Pes ON to resume monitoring. DOCUMENTATION AND MONITORING: Document in the Critical Care Respiratory Therapy Flowsheet the position at which the esophageal catheter is secured. Check the position a minimum of once per shift and PRN. Document transpulmonary pressures as follows: For patients on mandatory modes of ventilation, document the following values in the Critical Care Respiratory Therapy Flowsheet with each monitoring and with every parameter change: Ptp Plat Ptp PEEP Pes The values for each of the parameters displayed on the AVEA screen should be the value documented on the flowsheet UNLESS the waveform display indicates a value that is significantly different (> 3 - 5 cmH2O). In these cases, determine the potential reasons for the discrepancy and document the rationale for which value is being used in the narrative of the flowsheet. Document rationale for all changes made in the narrative of the flowsheet as per RTD5029 Ventilatory Monitoring Protocol. SPECIAL CONSIDERATIONS: The AVEA ventilator automatically cycles the balloon (deflates and re-inflates) to ensure that the measurements are accurate and not affected by leaks. During this time, the scalar will turn green and waveforms will be temporarily lost. For Maquet/Getinge Servo U with EPM software Equipment: Servo U Y Sensor Module Getinge Stop Cock adaptor or equivalent 4 way stop cock Artline extension tubing Luer compatible Esophageal Balloon 3 cc Syringe Blue pad Water Soluble Lubricant PROCEDURE FOR PREPARATION OF Servo U VENTILATOR: 1: Remove blank cartridge from Servo U 2: Insert Y sensor module into newly freed up slot, note that grey cable is not required 3: Ensure that Getinge ___ adaptor or equivalent is attached to Y sensor modules nipple 4: Servo U will automatically recognize and configure for EPM once module is inserted 5: Attach artline extension tubing ( __ inch) with 4 way stopcock PROCEDURE FOR PLACEMENT OF THE ESOPHAGEAL CATHETER: A: Cooper Medical Esophageal Balloon Pre use check and patient measurement Measure patient for NEX ( nose/ear/xyphoid) target placement will be lower 1/3rd of esophagus with patient supine and HOB at 30 degrees. Average female will likely be 35-40 cm. Average Male will be 40-45cm depth of insertion. Inflate Cooper balloon with 3 cc of air, ensure balloon is leak free. Withdraw the 3 cc of air once patency is confirmed. Place blue pad or equivalent over patients chest Lubricate balloon with water soluble lubricant Ensure no contra-indications exist a. Conventional insertion With Y sensor module open to atmosphere ensure that Pes measures 0cmh20 Consider freezing or soaking cooper cath in ice bath to increase catheter ridgidity Ensure water soluble lubricant is present Insert either nasally or orally, gently insert in 1-2 cm increments following hard palate if nasally, Insert to target depth ( we need to make a decision here, some say go to stomach level, inflate and withdraw till cardiac osc), attach 4 way stopcock and artline extension tube to Y sensor module Inflate cooper balloon with 1 cc of air, open 4 way stopcock to balloon and Y sensor. Verify cardiac oscillations present on Pes waveform. Note in real world generally not visible secondary to autoscaling, do the placement maneuver with out compression, wait the ten seconds look at the placement summary for cardiac osc If not present on Servo U screen with insertion of module, press Maneuvers(Press Pes Catheter Positioning( Follow Pes catheter position instructions and proceed with 1-2 gentle chest compression over sternum. This will verify the positive pressure occlusion ratio is within acceptable range of 0.8 to 1.2. The Servo U is incapable of telling if the catheter is high or low, however a ΔPes/ΔPaw ratio of > 1.2 indicates the balloon is to deep, a ΔPes/ΔPaw ratio of < 0.8 indicates the balloon is sitting to high. Optional If outside of range of 0.8 to 1.2 consider an expiratory hold with 1-3 gentle abdominal pumps with the record feature activated. A positive deflection on the Pl ee will indicate catheter not deep enough. A negative deflection on the Pl ee indicates catheter is too deep. No or minimal deflection on the PL ee with cardiac oscillation present indicates correct position. No or minimal deflection with no cardiac oscillation has a very likely chance that catheter has been inserted into lung ( one of the reasons I do not like inserting into gut and withdrawing) Adjust for depth as needed and repeat the Servo U catheter position maneuver to ensure proper placement. Once placement is verified measurements can be obtained. Pl ei, PL ee, PL drive, Change Pes. The Servo U will continuously measure these measurements. Performing Inspiratory Hold and Expiratory Hold for true measurement to ensure any air trapping and proper equilibration of pressures are transmitted to the ventilator. Verify with abdominal Xray, Cooper medical stainless steel guidewire should be visible on Xray. Once verified with X Ray for positioning, remove the stainless steel guidewire, reattach 4 way stopcock. Never reinsert the stainless steel guidewire due to perforation of catheter and patient tissues. Document in CST, insertion procedure, depth of catheter, orifice of insertion, placement verification, and relevant pulmonary mechanics numbers. b. Sima Technique insertion As the Cooper catheter tends to be very soft and floppy even with stainless guidewire an alternate technique may be required For insertion use a #18 french dual lumen NG, insert tip of Cooper catheter 1 cm into NG tube distal Murphy’s Eye. Ensure lubrication is present. Insert the NG with Cooper combination, NG will pull the Cooper down and into the patient. Stop inserting at desired depth. Follow inflation and measurement techniques as described in a, above. If in desired place, firmly hold the Cooper Catheter, continue to insert the NG tube another 10 to 15 cm. Cooper catheter should disengage from NG. Withdraw the NG tube from patient. B: Flux Med Respiratory Mechanics Kit ( Pending C: Nutrivent Esophageal Balloon ( Pending PROCEDURE FOR OBTAINING PRESSURE MEASUREMENTS: The Servo U Y sensor module is “hot swappable” and can transfer from ventilator to ventilator for measuring assuming proper sanitation has been followed. If balloon catheter is present, prior to measurement A: Verify depth of catheter B: Open 4 way stopcock to patient, use 3 cc syringe and withdraw all volume from balloon C: Open 4 way stopcock to atmosphere, ensure Pes is at zero D: Close stopcock to Atmosphere, inflate balloon with 1 cc of air using the 3 cc syringe E: Ensure cardiac oscillation present on Pes tracing F: To verify that catheter remains in correct position, Press maneuvers and follow the Servo U catheter positioning process as described above. G: Obtain measurements as described above. Patient Transport with Cooper Medical Catheter Simply disconnect extension tubing from 4 way stopcock, and verify on Servo U. ( need to reword, verify what the vent says) Recommendations Pl ei ( same as Ptp plat) should be kept less than 20-25 cmh20 Pl drive should be maintained < 10-12 cmh20 Pl ee ( same as Ptp peep) should be in target range of 0-5 cmh20 as per peep/fio2 table. REFERENCES: Talmor, D et al. Mechanical ventilation guided by esophageal pressure in acute lung injury. NEJM 2008; 359: 2095-2104. Talmor, D et al. Mechanical ventilation guided by esophageal pressure in acute lung injury/ARDS. NEJM Journal Club; Nov. 2008. Talmor, D et al. Esophageal and transpulmonary pressures in acute respiratory failure. Crit Care Med 2006; 34: 1389-1394. Gattinoni, L et al. Bench to bedside review: chest wall elastance in acute lung injury/acute respiratory distress syndrome patients. Critical Care 2004; 8: 350-355. |